# Risk–Benefit Analysis of Introducing SmartSweep in Antenatal Care

# 1. Overview

SmartSweep is a handheld ultrasound-based device designed to assist antenatal care professionals—such as midwives—in the evaluation of five key parameters:
- Detection of multiple pregnancies
- Identification of non-cephalic fetal lie
- Measurement of deepest vertical pocket (DVP) of amniotic fluid
- Verification of cardiac activity
- Estimation of gestational age

The current standard of care relies on manual methods:
- Palpation to identify multiple pregnancies and fetal lie
- Measurement of fundal height or LMP date for amniotic fluid or gestational age
- Auscultation for fetal heartbeat

Manual palpation is known to have low diagnostic accuracy; for instance, its sensitivity for detecting fetal position can be as low as ~34%, with a specificity of ~71%.

# 2. Baseline Risk of Current Standard of Care

| Parameter | Sensitivity (Approx.) | Specificity (Approx.) | False Negative Rate | False Positive Rate |
| --- | --- | --- | --- | --- |
| Multiple pregnancies | 60% | 60% | 40% | 40% |
| Non-cephalic lie | 34% | 71% | 66% | 29% |
| Gestational age (LMP) | Varies | — | — | — |
| Amniotic fluid (DVP) | Varies with skill | — | — | — |
| Cardiac activity | Limited early | — | — | — |

# 3. Probability Analysis and Failure Rate Margin (FRM)

We define a Failure Rate Margin (FRM) to quantify the statistical benefit of SmartSweep relative to current methods:

FRM = P_failure_manual – P_failure_SmartSweep

Where:

- P_failure_manual: Failure probability of the manual method (either false negative or false positive).

- P_failure_SmartSweep: Failure probability of SmartSweep, considering both user errors and device limitations.

SmartSweep Failure Probability:

P_failure = P_uf × P_ab_pred × P_prev_ab + (1 − P_uf) × P_ss_failure

Definitions:

- P_uf: Probability of User Failure (e.g., incorrect positioning or protocol)

- P_ab_pred: Condition prediction probability if user error occurs

- P_prev_ab: Prevalence of the condition in the population

- P_ss_failure: Device’s failure rate under proper usage

# 4. Example Simulations of FRM: Twins

Using the prediction of multiple pregnancies as an example, we can make a few assumptions based on our understanding of the system behavior:

- P_ab_pred (Failure probability if user error occurs) = 0.5. If the user fails to perform the protocol correctly, the parameter will be estimated at random.

- P_prev_ab (Prevalence of the condition in the population) = 0.98. 98% of pregnancies are singleton.

- P_ss_failure (Device’s failure rate under proper usage) = 0.1. SmartSweep is known to predict multiple pregnancies correct around 90% of the time if the user performs a proper exam.

Making these assumptions, we can simulate the effect of the likelihood of user failure P_uf on the FRM for the False Positive Rate of predicting multiple pregnancies

- If there is no chance of user failure, then the Failure Rate Margin (FMR) is positive (0.3). That is, the probability of a failure from the manual method—manual palpation, false positive rate, 0.4—is higher than the anticipated FPR of SmartSweep (0.1).

- If the user fails every time (P_uf = 1), then there is more benefit to the manual method. In this case, SmartSweep will randomly predict singleton or multiples, since P_ab_pred = 0.5.

- If the user fails occasionally (e.g. P_uf = 0.2), then the margin of benefit of SmartSweep starts to erode. The FRM drops to roughly 0.25, which means that SmartSweep, even with the user error rate of 20% will still statistically predict the right number of fetuses 25% more often than the manual method.

# 5. Assumptions Table

# 6. Assumptions Table

|  | Probability of Failure if Poor Input | Prevalence of Abnormal Condition | Probability of Failure if Sufficient Input | Probability of Failure with Manual Method | Notes |
| --- | --- | --- | --- | --- | --- |
| Twins, False Negative | 0.5 | 0.02 | 0.1 | 0.4 |  |
| Twins, False Positive | 0.5 | 0.98 | 0.1 | 0.4 |  |
| Lie, False Negative | 0.5 | 0.04 | 0.042 | 0.30 |  |
| Lie, False Positive | 0.5 | 0.96 | 0.01 | 0.40 |  |
| DVP_Oligo, False Negative | 0.5 |  |  |  |  |
| DVP_Oligo, False Positive | 0.5 |  |  |  |  |
| DVP_Poly, False Negative | 0.5 |  |  |  |  |
| DVP_Poly, False Positive | 0.5 |  |  |  |  |

# 7. Discussion

SmartSweep significantly lowers failure risk across scenarios (FRM ~35–36%), driven primarily by automated detection and lower intrinsic error. FRM is relatively robust to user error rates in this example. Actual values of P_uf, P_ab_pred, P_ss_failure, and P_prev_ab should be determined empirically for accuracy. Implementation success hinges on training to minimize P_uf and device calibration.

# 8. Next Steps

- tbd